## Slide 1
Evaluation of traffic’s influence near school front doors with low-cost PM2.5 monitoring C.Falzone (1), F.Simon (1), F.Lenartz (2), A-C.Romain (1)
Sensing of Atmospheres and Monitoring, University of Liege, Arlon, Belgium (2) ISSeP, Liege, Belgium
cfalzone@uliege.be
Sensing of Atmospheres and Monitoring
Objectives
Field tests comparison: low-cost sensors collocated with Air Quality monitoring station
Monitoring traffic emissions of PM2.5 during the children drop-off at school
Highlighting children’s exposure to PM2.5 during school period
Material and methods
Results and discussion
Field Test Comparison
3 lab-made devices, battery powered, (called 3PM, each one consisting of 3 low-cost sensors PM2.5 Honeywell HPM Series–HPMA115S0–XXX ;
0 - 1000 µg/m³ ; autonomy 9h)
GRIMM EDM180 as reference for calibration
3 days comparison at ISSeP, Liege (Belgium)
        Sensor vs Sensor
        Sensor vs GRIMM
  Sensor uncertainty
Conclusions
Reproducibility between sensors is very good
Expanded relative uncertainty for limit value (25 µg/m³ ; Reference Measure uncertainty = 0.67 µg/m³)**:
    No corrected sensor: 60% and Corrected sensor: 54%
Upcoming Technical Specification from the CEN/TC/264/WG42 “Air quality – Performance evaluation of air quality sensors” will be useful
First results highlight that the children drop-off to school has an impact on PM2.5 values
Additional study is required, involving various long term measurement campaigns and different schools
PM2.5 monitoring around the school
  Traffic profiles highlight peaks during the beginning, middle and end of school-hours
3PM correction* by Deming regression (orthogonal)
x = reference and y = sensor)
  Concentrations around the school are always higher than background concentrations
  Concentration profile follows the traffic profile
  PM2.5 concentration exceeds the limit value when the background concentration is approximatively 20 µg/m³
30 minutes average
5 minutes average
1 hour average
25 µg/m³
07:30
07:30
WeBIOPATR, 2019; Belgrade
PM2.5 monitoring around of the school
3PM fixed above each entrance of the school during 14 days including school days, week-ends and holidays. Traffic is measured in 3 streets surrounding school with a TMSSA radar
Rural school located in the south of Belgium
  n = measurment number ;
  p = replicas number ;
  yi,j = sensor measure on i period
  ym = mean of y on i period ;
Field Test Comparison
References:
*Equivalence test: BV/AQ/AD202209/DH/2396; **Directive 2008/50/CE and https://ec.europa.eu/environment/air/quality/legislation/pdf/RIVM_PM_equivalence_v2.9.xls
n=44 | a | u(a) | b | u(b) | R² | RMSE
3PM_1 | 1.225 | 0.036 | -7.966 | 1.622 | 0.964 | 4.36
3PM_2 | 1.335 | 0.045 | -11.874 | 2.07 | 0.951 | 4.25
3PM_3 | 1.348 | 0.045 | -11.247 | 2.042 | 0.953 | 5.49