## Slide 1
Verification of the reference intervals proposed by Abbott on Alinity ci in a population of healthy subjects from Liege, Belgium.
C.Bertholet1, M.Lamtiri-Laarif1, L.Vranken1 , R.Gadisseur1 , S.Delcour1 , E.Cavalier1
Introduction:
1 Department of Clinical Chemistry, University of Liège, CHU Sart-Tilman, B-4000 Liège, Belgium.
The institute for clinical chemistry, laboratory medicine and endocrinology at the university hospital (CHU)in Liege has introduced the Abbott Alinity ci system family in the routine testing.
Before routine start with the Alinity solution validation and verification of the new routine systems and the respective assays is important. Before CHU Liege was working with a Roche solution.
According to ISO 15189 standard, the analytical performances of CE-marked devices as provided by manufacturer in the respective instructions for use of the assays must be verified by certified laboratories. For each assay analytical performance (trueness and precision), limit of detection (LOD) and quantification (LOQ), measurement uncertainty and bias against peers and reference methods must be verified. Furthermore, a comparison of the new method to the previous solution has to be performed including samples with values distributed over the linear range of the assay..
Verification of the reference intervals as provided in the manufacturer instruction for use is also a key requirement.. Based on CLSI EP28-A3 guidelines samples from a selected population of 20 healthy individuals are required to verify the reference range. This is the preferred method as establishment of reference range for all analytes can be cumbersome. Depending on the analyte, the established reference interval can be influenced by age, sex, hormonal cycle or nychthemeral cycle. and a samples from 120 healthy individuals are required.
Objective of our evaluation was to verify the reference ranges of the Abbott Alinity clinical chemistry and immunoassay assay provided in the instructions for use using 80 samples coming from a population of apparently healthy Belgian individuals.
Material and Methods:
Materials: In total 80 healthy volunteers agreed to participate I the evaluation. In total 20 men and women below the age of 50 and 20 men and women above the age of 50 were enrolled. Blood was drawn after overnight fasting. Age of the population ranged between 24 and 70 years. We verified the reference intervals for basic chemistry (ions, liver enzymes, metabolites, proteins and lipids) and a wide panel of immunoassays (cardiac markers, tumor markers, fertility hormones, thyroid hormones and anemia panel). For fertility hormones, we checked references values for three groups: men, postmenopausal and premenopausal women. We used EP-Evaluator software following the CLSI EP28-A3 guidelines. The reference values as provided the manufacturer instruction for use were accepted if a maximum of 10% of healthy patients were outside the range with 95% IC.
Statistical analysis:
EP-Evaluator software was used to perform the verification of reference intervals.
Conclusions:
Our results confirm the reference values ​as provided in the Abbott instructions for use for all parameters, except Cholinesterase. Indeed, 30% of healthy men were outside the range. As a consequence reference range need to established by running at least 120 samples coming from healthy subjects.
Future evaluations are planned. For FSH, LH and Prolactin assays, samples from healthy women on contraception were found outside the range. The three hormonal parameter of the outsiders were consistent with each other. As women have hormonal cycle it’s difficult to develop reference values following the cycle. For estradiol, plan is to also run assays on mass spectrometry especially for children and men to hopefully achieve better accuracy in the lo range (LOQ = 23,7 pg/ml).
Email: c.bertholet@chuliege.be
Results
For basic chemistry, all results of the healthy volunteers were in accordance with the reference values provided in the instruction for use of the Abbott Alinity assays. Maximum number of results outside the proposed reference range was 10% percent.
With the exception of Cholinesterase results outside the proposed reference interval were below 10%.
For Cholinesterase 5% of healthy women were outside the range (2,88 – 12,67 kU/l) however 30% of healthy men had values that were outside(4,93 – 10,93 kU/l).
All results are summarized in Table 1.
For the Abbott Alinity immunoassays, all values of the healthy individuals were in agreement with Abbott’s proposed values.
Analyte | Reference Range -Alinity | CHU Reference Range | Assay Group
Amylase (U/L) | 25 – 125 | 25,64 – 85,95 | Basic Chemistry
Bicarbonate (mmol/L) | 22 – 31 | 23,09 – 34,86 (9,1 % rejected) | Basic Chemistry
Bilirubin D(mg/dL) | 0 – 0,5 | 0,13 – 0,51 (5% rejected) | Basic Chemistry
Bilirubin T (mg/dL) | 0,3 – 1,2 | 0,27 – 1,41 (10% rejected) | Basic Chemistry
Calcium (mmol/L) | 2,2 – 2,6 | 1,99 – 2,48 (10% rejected) | Basic Chemistry
Chloride (mmol/L) | 98 – 107 | 98,98 –114,62 (9,5% rejected) | Basic Chemistry
CRP (mg/L) | 0 – 5 | 0,13 – 3,19 | Basic Chemistry
Glucose (mg/dL) | 60 – 100 | 75,59 – 107,36 (10% rejected) | Basic Chemistry
GGT (U/L) | 12 – 64 | 12,25 – 139,94 (10% rejected) | Basic Chemistry
Potassium (mmol/L) | 3,5 – 5,1 | 2,97 – 4,43 (10% rejected) | Basic Chemistry
Lactate (mg/L) | 45 – 198 | 23,23 – 163,1 (10% rejected) | Basic Chemistry
LDH (U/L) | 125 – 220 | 112,56 – 282,9 (10% rejected) | Basic Chemistry
Magnesium (mmol/L) | 0,66 – 1,07 | 0,67 – 0,96 | Basic Chemistry
Sodium (mmol/L) | 136 – 145 | 129,2 – 143,94 (10% rejected) | Basic Chemistry
Phosphates (mmol/L) | 0,74 – 1,52 | 0,71 – 1,27 (5% rejected) | Basic Chemistry
Prealbumine (g/L) | 0,16 – 0,45 | 0,18 – 0,4 | Basic Chemistry
Total proteins (g/L) | 58 – 83 | 49,32 – 80,21 (5% rejected) | Basic Chemistry
Triglycerides (mg/dL) | 0 – 150 | 3,05 – 199,27 (10% rejected) | Basic Chemistry
TGO (U/L) | 5 – 34 | 9,47 – 63,8 (5% rejected) | Basic Chemistry
TGP (U/L) | 0 – 55 | 7,82 – 72,02 (5% rejected) | Basic Chemistry
Lipase(U/L) | 0 – 60 | 13,2 – 69,4 (5% rejected) | Basic Chemistry
SB2 (mg/L) | 0,97 – 2,64 | 1,282 – 2,338 | Basic Chemistry
Uric acid (mg/dL) | Men : 3,5 – 7,2 | 4,22 – 9,12 (5% rejected) | Basic Chemistry
 | Women : 2,6 – 6 | 3,26 – 6,54 (5% rejected) | Basic Chemistry
C3 (g/L) | Men : 0,82 – 1,85 | 0,72 – 1,53 (5% rejected) | Basic Chemistry
 | Women : 0,83 – 1,93 | 0,81 – 1,68 (5% rejected) | Basic Chemistry
C4 (g/L) | Men : 0,15 – 0,53 | 0,15 – 0,35 | Basic Chemistry
 | Women : 0,15 – 0,57 | 0,14 – 0,46 (10% rejected) | Basic Chemistry
Cholinesterase (kU/L) | Men : 4,39 – 10,93 | 5,68 – 15,25 (30% rejected) | Basic Chemistry
 | Women : 2,88 – 12,67 | 5,33 – 13,66 (5% rejected) | Basic Chemistry
Creatinine kinase | Men : 30 – 200 | 39,22 – 2897,81 (10% rejected) | Basic Chemistry
(U/L) | Women : 29 - 168 | 35,63 – 135,3 | Basic Chemistry
Creatinine (mg/dL) | Men : 0,73 – 1,18 | 0,73 – 1,31 (5% rejected) | Basic Chemistry
 | Women : 0,55 – 1,02 | 0,56 – 0,92 | Basic Chemistry
Iron (µmol/L) | Men : 11,6 – 31,3 | 11,43 – 24,0 3(5% rejected) | Basic Chemistry
 | Women : 9 – 30,4 | 8,93 – 30,02 (5% rejected) | Basic Chemistry
IgA(g/L) | Men : 0,63 – 4,84 | 0,05 – 3,54 (5% rejected) | Basic Chemistry
 | Women : 0,65 – 4,21 | 1,1 – 3,47 | Basic Chemistry
IgG(g/L) | Men : 5,4 – 18,22 | 5,67 – 16,13 | Basic Chemistry
 | Women : 5,52 – 16,31 | 6,71 – 12,82 | Basic Chemistry
IgM(g/L) | Men : 0,22 – 2,4 | 0,25 – 2,34 | Basic Chemistry
 | Women : 0,33 – 2,93 | 0,6 – 2,13 | Basic Chemistry
Alkaline phosphatase (U/L) | Men : 40 – 150 | 38,11 – 101,5 (5% rejected) | Basic Chemistry
 | Women : 40 - 150 | 34,24 – 83,25 (10% rejected) | Basic Chemistry
Transferrine (g/L) | Men : 1,74 – 3,64 | 3,96 – 61,6 (10% rejected) | Basic Chemistry
 | Women : 14,952 – 39,94 | 13,4 – 35,29 (5% rejected) | Basic Chemistry
Albumin (g/l) | 20 – 60 yo : 35 – 52 | 35,72 – 51,15 | Basic Chemistry
 | > 60 yo : 32 – 46 | 29,79 – 46,04 (10% rejected) | Basic Chemistry
Urea (mg/dl) | W < 50 yo : 14,95 – 39,94 | 13,4 – 35,29 (5% rejected) | Basic Chemistry
 | W > 50 yo : 20,93 – 42,93 | 21,46 – 38,48 | Basic Chemistry
 | M < 50 yo : 19,01 – 44 | 18,82 – 40,5 | Basic Chemistry
 | M > 50 yo : 17,94 – 54,89 | 23,63 – 49,55 | Basic Chemistry
Free T3 (pmol/l) | 2,8879 – 4,8848 | 2,93 – 4,81 | Thyroid
Free T4 (pmol/l) | 9,01 – 19,05 | 8,91 – 14,92 (5% rejected) | Thyroid
TSH (mUI/l) | 0,35 – 4,94 | 0,114 – 3,5473 (5% rejected) | Thyroid
Anti-TPO (IU/ml) | 0 – 5,61 | 0,01 – 1,95 | Thyroid
Anti-TG (IU/ml) | 0 – 4,11 | 0,8 – 4,49 (10% rejected) | Thyroid
Folates (ng/ml) | 3,1 – 20,5 | 3,18 – 13,41 | Anemia and other panel
B12 (pg/ml) | 187 – 883 | 168,65 – 518,03 (5% rejected) | Anemia and other panel
Ferritin (µg/l) | Men : 21,81 – 274,66 | 23,04 – 309,69 (10% rejected) | Anemia and other panel
 | Women : 4,63 – 204 | 20,24 – 350,52 (5% rejected) | Anemia and other panel
Insulin (pmol/l) | 0 – 104 | 10,02 – 77,48 | Anemia and other panel
C peptid (nmol/l) | 0,259 – 1,729 | 0,3067 – 0,8957 | Anemia and other panel
Cortisol (nmol/l) | AM : 102,1 – 535,2 | 191,83 – 546,12 (5% rejected) | Anemia and other panel
 | PM : 80 – 477,3 | 70,89 – 219,28 (10% rejected) | Anemia and other panel
HCG (IU/l) | 0 – 5 | 0 - 4 | Fertility Hormones
Prolactin (µg/l) | Men : 3,46 – 19,4 | 3,96 – 61,6 (10% rejected) | Fertility Hormones
 | Women :,5,18 – 26,53 | 4,44 – 31,25 (10% rejected) | Fertility Hormones
SHBG (nmol/l) | Men : 13,5 – 71,4 | 20,31 – 53,79 | Fertility Hormones
 | Women : 19,8 – 155,2 | 24,1 – 238,29 (10% rejected) | Fertility Hormones
Testosteron (nmol/l) | W < 50 yo : 0,48 – 1,85 | 0,45 – 1,44 (5% rejected) | Fertility Hormones
 | W > 50 yo : 0,43 – 1,24 | 0,53 – 1,38 (10% rejected) | Fertility Hormones
 | M < 50 yo : 8,33 – 30,19 | 8,21 – 24,08 (5% rejected) | Fertility Hormones
 | M > 50 yo : 7,66 – 24,82 | 7,45 – 24,82 (5% rejected) | Fertility Hormones
LH (IU/l) | premenopausal women : 0,56 –14 | 0,02 – 15,72 0,03 – 10,48 (25% rejected) | Fertility Hormones
 | Postmenopausal women : 5,16 – 61,99 | 0,01 – 44,61 (5% rejected) | Fertility Hormones
 | Men : 0,57 – 12,07 | 1,41 – 11,36 | Fertility Hormones
FSH (IU/l) | premenopausal women : 1,38 – 8,08 | 0,03 – 10,48 (25% rejected) | Fertility Hormones
 | Postmenopausal women : 26,72 – 133,41 | 0,19 – 104,13 (10% rejected) | Fertility Hormones
 | Men : 0,95 – 11,95 | 1,9 – 42,67 (5% rejected) | Fertility Hormones
Estradiol (pg/ml) | premenopausal women : 21 – 312 | 0 – 219,67 (40% rejected) | Fertility Hormones
 | Postmenopausal women : 0 – 28 | 0 – 106,44 (15% rejected) | Fertility Hormones
 | Men : 11 - 44 | 9,49 – 37,59 (10% rejected) | Fertility Hormones
Progesterone (ng/mL) | premenopausal women : 0,1 – 0,3 | 0 – 11,42 (60% rejected) | Fertility Hormones
 | Postmenopausal women : 0 – 0,2 | 0 – 1,86 (5% rejected) | Fertility Hormones
 | Men : 0 – 0,5 | 0 – 0,32 | Fertility Hormones
AFP (ng/ml) | 0,89 - 8,78 | 1,42 - 6.24 | Tumor Marker
CA 15-3 (U/ml) | 0 - 31,3 | 4,06 - 22,81 | Tumor Marker
CA 19-9 (U/ml) | 0 -37 | 0,03 - 14,84 | Tumor Marker
CEA (ng/ml) | 0 - 5 | 0,65 - 4,26 | Tumor Marker
TPSA (µg/l) | 0 - 4 | 0 - 3.276 | Tumor Marker
Table 1
CHU Reference range and Alinity Reference Range for the different assays/panels