## Slide 1: FIRST DEVELOPMENT OF A GAS SENSOR ARRAY FOR AMMONIA EMISSION MONITORING FROM GRASSLANDS
Marie SCHEUREN; Anne-Claude ROMAIN
Sensing of Atmospheres and Monitoring (SAM-Lab), Department of Environmental Sciences and Management,
Research Unit SPHERES
University of Liege, - Avenue de Longwy 185, 6700 Arlon (Belgium). mscheuren@uliege.be
9th IWA Odour & VOC/Air Emission Conference
26-27 October 2021 Bilbao, Spain

## Slide 2: Introduction
Agriculture is the main source of ammonia emissions
  93% of total EU emissions
  Due to soil fertilisation and livestock manure storage
This study focuses on ammonia emissions caused by nitrogen fertilisation of grassland
Objective
  Development of a metal oxide sensor array for real time ammonia concentrations and odour measurement
    Dynamic olfactometry to measure odour
    Acid-soaked filters to determine NH3 concentration

## Slide 3: Materials and methods
Acid soaked filters
Material
  Filters are placed in a custom-made Teflon sealed chamber
  +12% solution of citric acid in methanol
Efficiency tests of the filters in the lab
  Several ammonia concentrations (gas cylinders)
  Various humidity values
  (Dilution unit consisting of mass flow controllers (MFC))
The filters are exposed to ambient air 60 minutes
during outdoor experiments

## Slide 4: Materials and methods
Odour sample
  Dynamic olfactometry according to the European standard EN 13725
  Olfasense TO evolution olfactometer
  4 panelists
  Max. 2 hours storage
Sampling every 3 hours
during the first 6 hours after fertilizer spreading and
at 9:00 the second and the third day

## Slide 5: Materials and methods
Metal oxide sensor array
  Sensors selected for their sensitivity to ammonia
  Placed in a hermetic PTFE chamber with temperature controlled at 50°C
  Temperature and humidity monitoring
  Ambient air sucks through the chamber at 200 ml/min
  (2 devices)
Sensors | Compound specified by the manufacturer | Concentration (ppm) specified by the manufacturer
TGS 2602 | NH3, H2S, COV | 1-30
TGS826 | NH3 | 30-300
SB-53-01 | NH3 | -
GGS4330 | NH3 | -
TGS 2620 | COV | 50-5000
GGS 2330 | CO, H2, C2H5OH | 1-1000

## Slide 6: Material and methods
Outdoor experiments
  50 m2 grassland
  Fertilisation of 100 nitrogen units of cattle slurry
  Air intake of all samples at a height of 30 cm
  June 2021
The study site is equipped with a weather station

## Slide 7: Results Ammonia concentration
Efficiency tests of the filters
  Efficient trapping of ammonia in the air (in the lab)
    R2 = 0.99
Filters: Outdoor experiment
  Before fertilisation ~10 ppb
  After fertilisation
    2 first hours ~1 ppm
    Succeeding days, the concentration decreased below 100 ppb with a peak around 12 o’clock
~ 12 o’clock

## Slide 8
Sampling (and measurement) day | Odour (uoE/m³) | Standard deviation | Ammonia (ppb) | Time after fertilisation
08/06/2021 09:00 | 572 | 333 | 1158 | Few minutes
08/06/2021 12:00 | 641 | 390 | 462 | 3 hours
08/06/2021 15:00 | 564 | 231 | 302 | 6 hours
09/06/2021 09:00 | 212 | 234 | 83 | 24 hours
10/06/2021 09:00 | 232 | 216 | 26 | 48 hours
Results Odour concentration
  First odour concentration (just after spreading) too low regarding the high NH3 value (1158 ppbv);
  Except for very high value of NH3, a trend between odour and NH3 is observed
   NH3 is not the main contributor of odour but could be a good indicator of spreading odour when NH3 is not too high
Chart: Odour (uoE/m³)
| 572 | 641 | 564 | 212 | 232
 | 1158.44 | 462 | 302.24 | 82.8 | 26.18

## Slide 9: Results Metal oxide sensor array response
Outdoor experiment
  Day 1
    Resistance (R) variation due to the fertilisation
    R increase in accordance with the emission decrease
  Day 2
    Slightly signal decreases
  Day 3
    Signal seems similar to days before fertilisation
Fertilisation

## Slide 10: Results Metal oxide sensor array response
  Day 1: 69 obs. 6 variables (value every 5 minutes)
    First 6 hours after fertilisation
PCA analysis (raw R values in kohms)
  Campaign : 193 obs. 6 variables (hourly median)
    3 days before + 6 days after
    PC 1 (Dim1) = time evolution of the signal
  PC 1 (Dim1) = air composition
  PC 2 (Dim2) = impact of humidity

## Slide 11: Conclusions
    Ammonia concentration
      Day 1: first two hours, NH3 increases to ~1 ppm, fast decrease to low level (<500 ppb)
      Day 2: value of ~ 100 ppb
      Day 3: level similar to background (<20 ppb)
      Peak at about 12 o'clock on the second and third day after fertilisation = temperature and sunshine
Odour concentration
      Day 1: High concentration with low increase at 12 o'clock
      Day 2 and 3: lower as day 1
Metal oxide sensor array
  Correlation with NH3 concentration
  No direct link with the odour concentration but with the chemical composition
Acid soaked filters = good tool for ammonia concentration validation
Metal oxide sensor array = promising tool for low cost ammonia real time monitoring
Next step
Repeat the tests and increase the dataset
  Quantification algorithm development from sensors signals

## Slide 12: Thank you for your kind attention

## Slide 13: Acknowledgements to
Intercomparison campaign with various analysers