Acetylcysteine and enzymatic creatinine: beware of laboratory artefact!

Michael Lognard1, Etienne Cavalier1, Jean-Paul Chapelle1, Bernard Lambermont2, Jean-Marie Krzesinski3, Pierre Delanaye3

1CHU Sart Tilman, Department of Clinical Chemistry, Liège, Belgium

2CHU Sart Tilman, Department of Medicine, Medical Intensive Care Unit, Liège, Belgium

3 University of Liege, Service de Dialyse, CHU Sart Tilman, Department of Nephrology, Dialysis, and Hypertension, 4000 Liège, Belgium

Sir: N-Acetylcysteine (NAC) is an antioxidant agent used as a mucolytic and to prevent acute hepatic failure following paracetamol poisoning [1]. NAC has also been thought to prevent radiocontrast nephrotoxicity [2]. Potential analytical interferences between NAC and creatinine measurement have been poorly studied. Tepel et al. showed no interference between their Jaffé creatinine and NAC in vitro [2]. Enzymatic methods are more precise and specific for creatinine measurement than the classic Jaffe methods [3]. Enzymatic assays are thus recommended. However, creatinine enzymatic assays are not free from interference, as has been described with flucytosine and high doses of dobutamine [3].

We have studied, in vitro, potential interferences between NAC and creatinine measurement by Jaffe and enzymatic methods. We decided to study three serum pools with creatinine concentrations around 1, 2 and 4.5mg/dl. Each sample was analysed in duplicate before and directly after adding increasing doses of NAC from 500 to 10,000 µg/ml (ampoule at a concentration of 100,000 µg/ml), using both the Jaffe method (CREA Jaffé; Roche Modular, Mannheim, Germany) and the enzymatic method (CREA Plus; Roche Modular, Mannheim, Germany). Volumes of the samples were adjusted to avoid dilution interference. Interference was considered as significant when the difference between the values measured before and after NAC addition was over 10% [4]. We performed the experiment in triplicate. The mean creatinine values of pools were thus 1.01 ± 0.24 mg/dl, 2.14 ± 0.07 mg/dl and 4.83 ± 0.49 mg/dl.

According to our criterion, we did not observe any interference with the Jaffe method. However, significant interferences were found when the enzymatic method was used from an acetylcysteine concentration of 1,000 µg/ml upward. The interferences increased with acetylcysteine concentrations (mean differences for the three pools went from -3.7% for an acetylcysteine concentration of 500 µg/ml to 71.7% for an acetylcysteine concentration of 10,000 µg/ml) (Table 1). In the three pools of creatinine, we observed a strict linear decrease in creatinine values with increasing doses of NAC (between 500 and 7,500 µg/ml) (r = 0.99).

In this study, we show that NAC can interfere with the dosage of creatinine by enzymatic assay and lead to an underestimation of creatinine concentration. The enzymatic technique involves the liberation of hydrogen peroxide, which produces colour by oxidizing a chromogenic dye [3, 5]. As it has powerful antioxidant effect, NAC may thus theoretically interfere with the coloration. However, significant interference occurs only at high NAC concentration (over 1,000 µg/ml), whatever the initial serum creatinine concentration. As NAC is quickly metabolized by the organism (plasma concentration of 20 µg/ml at 2 h after injection of 600 mg), such analytical interference is of clinical importance only in specific circumstances. Even when NAC is administered at very high dosage in the case of paracetamol poisoning, the average plasma NAC concentration is 554 µg/ml after a loading dose of 150 mg/kg [1]. At such plasma concentrations, our in vitro results confirm the absence of significant laboratory artefact of NAC in the enzymatic creatinine measurement. However, if the blood sample is drawn on a multiple-lumen catheter where NAC is perfused and which has not been correctly clamped before sampling, high concentrations of NAC may be present in the sample. In this case, a significant falsely low value of creatinine might be obtained with the enzymatic method, as has been described for dobutamine [5].

Table 1: Interference of N-acetylcysteine (NAC) in enzymatic creatinine measurements (expressed in mg/dl)

| Concentration of NAC added in vitro (µg/ml) | Three pools around 1 mg/dl (mean: 1.01 mg/dl): mean differences in % from base-line concentrations (standard deviation) | Three pools around 2 mg/dl (mean: 2.14 mg/dl): mean differences in % from base-line concentrations (standard deviation) | Three pools around 4.5 mg/dl (mean: 4.83 mg/dl): mean differences in % from base-line concentrations (standard deviation) |
| --- | --- | --- | --- |
| 500 | -3.9% (±1.1%) | -4.7% (± 0.8%) | -2.4% (±1.1%) |
| 1,000 | -ll.l%(±2.1%)a | -9.6% (± 1.7%) | -7.5% (±1.2%) |
| 5,000 | -44.8%(±11.9%)a | -51.6%(±8.2%)a | -50% (± 11.4%)a |
| 7,500 | -63.6% (± 15.9%)a | -62.4% (± 10.7%)a | -62.6% (± 10.2%)a |
| 10,000 | -70.7(±11.2%)a | -70.3% (± 6.3%)a | -74.1%(±5.8%)a |

aDifference over 10%, viewed as clinically significant

References

1.    Prescott LF, Park J, Ballantyne A, Adri-aenssens P, Proudfoot AT (1977) Treatment of paracetamol (acetaminophen) poisoning with N-acetylcysteine. Lancet 2:432-434

2.    Tepel M, van der Giet M, Schwarzfeld C, Laufer U, Liermann D, Zidek W (2000) Prevention of radiographic -contrast-agent-induced reductions in renal function by acetylcysteine. N Engl J Med 343:180-184

3.    Perrone RD, Madias NE, Levey AS (1992) Serum creatinine as an index of renal function: new insights into old concepts. Clin Chem 38:1933-1953

4.    Sonntag O, Scholer A (2001) Drug interference in clinical chemistry: recommendation of drugs and their concentrations to be used in drug interference studies. Ann Clin Biochem 38:376-385

5.    Daly TM, Kempe KC, Scott MG (1996) "Bouncing" creatinine levels. N Engl J Med 334:1749-1750