Non-destructive measurement of the magnetic moment of bulk GdBCO samples using a torque magnetometer designed for large size samples

Sébastien Brialmont, Jean-François Fagnard and Philippe Vanderbemden

University of Liege, Department of Electrical Engineering & Computer Science B28, 4000 Liege, Belgium

The non-destructive determination of the volume magnetic properties of large bulk superconductors requires the development of measurement techniques applicable to large size samples. We report the design for a magnetic torque measurement system able to characterize the magnetic moment of ~1 cm3 samples at room and cryogenic temperature. This device can be used for torque magnetometry with an intermediate sensitivity between cryogenic torque magnetometers designed for millimetric-size samples and commercial torquemeters poorly adapted to extreme conditions. The torque sensing system is based on four piezoresistive strain gauges cemented on a cylindrical transmission shaft and mounted in a full Wheatstone bridge configuration. The absolute sensitivity of the device was calibrated using a small coil fed with DC current. The torque magnetometer was tested on several magnetic samples in order to measure their magnetic moment under crossed field conditions: a Nd-Fe-B permanent magnet and a GdBCO bulk superconductor. An excellent agreement was found with results obtained with a flux extraction magnetometer. The system can be used to determine the average Jc of large samples and to study the demagnetization effects resulting from crossed magnetic fields. The torque magnetometer is able to measure magnetic moments ranging from 10-3 Am² (1 emu) up to a few Am². This extends significantly the accessible measurement range of “off-the shelf” magnetometers (e.g. SQUID, VSM) and could be scaled up to larger samples.

Acknowledgments:

We greatly acknowledge Nippon Steel for providing bulk large grain GdBa2Cu3O7 (GdBCO) samples.