Analytical Validation of the Maglumi for Gastrin measurement: Performance and Comparison with radioimmunoassay method.

Le Goff Caroline, Ferrante Nunzio, Lukas Pierre, Huyghebaert Loreen, Massonnet Philippe,  Dubrowski Thomas, Peeters Stéphanie, Etienne Cavalier

Department of Clinical Chemistry, University of Liège, CHU Sart-Tilman, B-4000 Liège, Belgium.

Objectives:

Gastrin (GST), a peptide hormone responsible for stimulating gastric acid secretion, plays a crucial role in diagnosing gastrointestinal diseases. Elevated or suppressed GST levels are associated with conditions such as Zollinger-Ellison syndrome, characterized by gastrin-secreting tumors, and chronic atrophic gastritis, linked to pernicious anemia or Helicobacter pylori infections. Accurate GST measurement is critical for clinical diagnoses and treatment. This study validates the Maglumi X3 (Snibe) chemiluminescent immunoassay (CLIA) analyzer for GST quantification in human serum and compares its results with established laboratory methods.

Materials and Methods:
The Maglumi X3 analyzer quantified GST in human serum samples. Performance was evaluated at five concentration levels (2.2 – 321 pmol/L) for intra-assay and inter-assay precision, trueness, and measurement uncertainty, following European Medicines Agency (EMA) guidelines. The validation standards were analyzed in triplicate over three days. Method comparison used 80 residual patient samples to assess Maglumi’s performance against DiaSource radioimmunoassay (RIA). Statistical analyses included Passing-Bablok regression and Bland-Altman plots via MedCalc software, with analytical validations supported by Enoval software (Arlenda).

Results:
The Maglumi X3 analyzer showed intra-assay and inter-assay coefficient of variations (CVs) at 3.26% and 5.30% for GST. Trueness evaluation revealed a maximum relative bias of 6.00%. Measurement uncertainty analysis showed a maximum relative expanded uncertainty of 11.3%. The Passing-Bablok regression equation was: GSTMaglumi = -2.017 + 0.2233 GSTRIA (95% CI for intercept: -3.465 to -0.8477, 95% CI for slope: 0.1902 to 0.2572). Systematic and proportional differences were observed. Bland-Altman plots confirmed agreement, there was no impact on the clinical diagnostic.

Conclusion:
The Maglumi X3 analyzer provides accurate GST measurements with acceptable precision and uncertainty. Despite systematic and proportional differences compared to RIA, clinical diagnostic outcomes remain unaffected. New reference intervals for Maglumi X3 are recommended to enhance diagnostic accuracy. The device is a valuable tool for assessing GST-related disorders and has replaced our RIA assay.